Laboratory Collection Guidelines for QuantiFERON-TB Gold PLUS

Figure 1

One QFT-PLUS test uses the following 4 collection tubes

All four tubes must be collected to perform test (Figure 1):

- Tubes are obtained from the Laboratory.

Filling Tubes

- Tubes should be at Room Temperature at the time of blood filling.

- Collect 1 mL of blood directly into each QFT-PLUS blood collection

Figure 2. Approved fill volume, image for guidance only

tube using a syringe and needle.

- DO NOT use a butterfly needle for QFT-PLUS collection.

- The black mark on the side of the tubes indicates the 1 mL fill volume.

- If the level of blood in any tube is not close to the indicator line,

another blood sample must be collected.  (Figure 2)

- If any of the four tubes is under or overfilled when sent to the laboratory,

the entire specimen will be rejected and MUST be recollected

Mixing

- Antigens are dried onto the inner wall of the blood collection tubes.

- Thoroughly mix tube contents with the blood to dissolve the heparin,

prevent clotting, and resolubilize the stimulating antigen.

- Mixing is performed by shaking the tubes 10 times just firmly enough

Figure 3. A correctly mixed tube

to ensure that the entire inner surface of the tube is coated with blood.

- Frothing of blood in the QFT-PLUS blood collection tubes after the

shaking process is normal (Figure 3).

- Place patient label long way on tubes and DO NOT cover the window

of the tube or the black fill line.

Send to Laboratory Immediately for Incubation

- Tubes must stay at room temperature.  Do not refrigerate.

No testing will be performed if tubes are not filled properly

Questions should be directed to your hospital laboratory.